## Stadiometer QA
STADIOMETER QUALITY CONTROL |  | STADIOMETER SHOULD BE CHECKED ANNUALLY USING A VALIDATED MEASURING ROD
Range | 80 ± 0.05cm
79.9 - 80.1
Date
Range
Tester

## Equipment Maintenance Log
Equipment Maintenance Log |  |  | It is important that each spiromter in the practice is maintained according to the manufacturers instruction. Spirometry equipment requires calibration,maintenance, error reporting, software updates quaterly and maintenance when it is due | For video on quality assurance and how to perform a syringe slow leak test click on link | Link to RSHQ QA video
Please enter initials and date | Spirometer Make: | Model Number: | Maintenance Service Calibration Due: | Calibraton syringe annual service | Syringe leak test | Syringe slow leak test
Electrical cable checks completed (TO BE COMEPLTED MONTHLY) | Check complete on equipment. Check housing, sensors etc for cracks or damage. (TO BE COMPLETED WEEKLY) | Software updates (should be upchecked quaterley) | Notes/Error Details (software version can be entered here) | Completed by on date | Completed by on date | Completed by on date
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
Q1
Q2
Q3
Q4
Your syringe is required to be calibrated annually to ensure that it is working accurately. You can organise this by contacting Bird Healthcare on 1300365561

## Log BTPS ON
| SPIROMETRY QA LOG NON BIOLOGICAL CONTROL | [5 empty] | This sheet should be used for spirometers were BTPS cannot be turned off when performing non-biological controls
[6 empty] | Automatically calculates corrected value should be between 2.91 and 3.09 |  | Automatically calculates corrected value should be between 2.91 and 3.09 |  | Automatically calculates corrected value should be between 2.91 and 3.09
Date | Time | Technician | FVC Low Flow Check (PEF<2L/s) | Temperature | Correction Factor | FIVC Low Flow Check (PEF<2L/s) | FIVC Low Flow corrected | FiVC Medium Flow Check (PEF between 4-6L/s) | FIVC Medium Flow Corrected | FIVC High Flow Check(PEF>8l/s) | FIVC High Flow Corrected | Comments
2025-04-30 | 216:00 | Technician 1 |  | 19 | 1.107 | 3.32 | 2.9991 | 3.32 | 2.9991 | 3.32 | 2.9991
|  |  |  | 20 | 1.102 |  | 0 |  | 0 |  | 0
|  |  |  | 18 | 1.113 |  | 0 |  | 0 |  | 0
|  |  |  | 21 | 1.097 |  | 0 |  | 0 |  | 0
|  |  |  | 23 | 1.086 | 3.32 | 3.0571 |  | 0 |  | 0
|  |  |  | 23 | 1.086 | 3.32 | 3.0571 |  | 0 |  | 0
|  |  |  | 23 | 1.086 |  | 0 |  | 0 |  | 0
|  |  |  | 23 | 1.086 |  | 0 |  | 0 |  | 0
|  |  |  | 23 | 1.086 |  | 0 |  | 0 |  | 0
|  |  |  | 23 | 1.086 | 3.32 | 3.0571 |  | 0 |  | 0
|  |  |  | 23 | 1.086 |  | 0 |  | 0 |  | 0
|  |  |  | 23 | 1.086 |  | 0 |  | 0 |  | 0
|  |  |  | 23 | 1.086 |  | 0 |  | 0 |  | 0
|  |  |  | 25 | 1.074 | 3.32 | 3.0912 | 3.32 | 3.0912 | 3.32 | 3.0912

## Log BTPS off
| SPIROMETRY QA LOG NON BIOLOGICAL CONTROL |  |  |  |  | Please remember that The ATPS box must be selected to turn off the BTPS correction factor when the syringe is being used in patient mode.
Date | Time | Technician | FVC Low Flow Check (PEF<2L/s) | FIVC Low Flow Check (PEF<2L/s) | FVC Medium Flow Check (PEF<4-6L/s) | FiVC Medium Flow Check (PEF between 4-6L/s) | FVC High Flow Check(PEF>8l/s) | FIVC High Flow Check(PEF>8l/s)2 | Comments
2025-04-30 | 216:00 | Technician 1

## BTPS Correction Table
Temp | Value | Correction Factor | Corrected Value
18 | 3.32 | 1.113 | 2.98
19 | 3.32 | 1.107 | 3
20 | 3.32 | 1.102 | 3.01
21 | 3.32 | 1.097 | 3.03
22 | 3.32 | 1.091 | 3.04
23 | 3.32 | 1.086 | 3.06
24 | 3.32 | 1.08 | 3.07
25 | 3.32 | 1.074 | 3.09
26 | 3.32 | 1.069 | 3.11
27 | 3.32 | 1.063 | 3.12
28 | 3.32 | 1.057 | 3.14

## Lookup Technician
Select Technician
Technician 1
Technician 2
Technician 3
Technician 4